## Modbus
Stav rele / State of relays
0x01 | 0x0000 | 0x0008
Externi vstupy / External inputs
0x02 | 0x0000 | 0x0008
Identifikace / Indentification
0x11
Výkonné příkazy / Do action
0x05 | 0x0000 | 0xFF00 | 0 |  |  | Zeroed value on display.
0x05 | 0x0001 | 0xFF00 | 1 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0002 | 0xFF00 | 2 |  |  | Zeroed value on display.
0x05 | 0x0003 | 0xFF00 | 3 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0004 | 0xFF00 | 4 |  |  | Zeroed value on display.
0x05 | 0x0005 | 0xFF00 | 5 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0006 | 0xFF00 | 6 |  |  | Zeroed value on display.
0x05 | 0x0007 | 0xFF00 | 7 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0008 | 0xFF00 | 8 |  |  | Zeroed value on display.
0x05 | 0x0009 | 0xFF00 | 9 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x000A | 0xFF00 | 10 |  |  | Zeroed value on display.
0x05 | 0x000B | 0xFF00 | 11 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x000C | 0xFF00 | 12 |  |  | Zeroed value on display.
0x05 | 0x000D | 0xFF00 | 13 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x000E | 0xFF00 | 14 |  |  | Zeroed value on display.
0x05 | 0x000F | 0xFF00 | 15 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0010 | 0xFF00 | 16 |  |  | Zeroed value on display.
0x05 | 0x0011 | 0xFF00 | 17 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0012 | 0xFF00 | 18 |  |  | Minimum and maximum clear
0x05 | 0x0013 | 0xFF00 | 19 |  |  | Clear memory of measured datas
0x05 | 0x0014 | 0xFF00 | 20 |  |  | Zeroed value on display.
0x05 | 0x0015 | 0xFF00 | 21 |  |  | Cancel display resetting (Tare clear)
0x05 | 0x0016 | 0xFF00 | 22 |  |  | Setting password for access to 0
Vyčtení naměřených a vypočtených hodnot / Reading of measured or evaluated values
0x04 | 0x0004 | 0x0002 | 30004 | float | FIL. A
0x04 | 0x0008 | 0x0002 | 30008 | float | FIL. B
0x04 | 0x000C | 0x0002 | 30012 | float | FIL. C
0x04 | 0x0010 | 0x0002 | 30016 | float | FIL. D
0x04 | 0x0014 | 0x0002 | 30020 | float | FIL. E
0x04 | 0x0018 | 0x0002 | 30024 | float | FIL. F
0x04 | 0x001C | 0x0002 | 30028 | float | FIL. G
0x04 | 0x0020 | 0x0002 | 30032 | float | FIL. H
0x04 | 0x0022 | 0x0002 | 30034 | float | MAT. FN.
0x04 | 0x0024 | 0x0002 | 30036 | float | MIN.
0x04 | 0x0026 | 0x0002 | 30038 | float | MAX.
0x04 | 0x003C | 0x0002 | 30060 | float | TARE A
0x04 | 0x003E | 0x0002 | 30062 | float | TARE B
0x04 | 0x0040 | 0x0002 | 30064 | float | TARE C
0x04 | 0x0042 | 0x0002 | 30066 | float | TARE D
0x04 | 0x0044 | 0x0002 | 30068 | float | TARE E
0x04 | 0x0046 | 0x0002 | 30070 | float | TARE F
0x04 | 0x0048 | 0x0002 | 30072 | float | TARE G
0x04 | 0x004A | 0x0002 | 30074 | float | TARE H
0x04 | 0x004C | 0x0002 | 30076 | float | TAR.ACT
0x04 | 0x0060 | 0x0002 | 30096 | float | COLD. J.
Vyčtení hodnot z menu / Reading menu item values
0x03 | 0x0018 | 0x0001 | 40024 | List | INPUTS – CONFIG. – READ./S | Measuring rate | 0= 40.0, 1= 20.0, 2= 10.0, 3= 5.0, 4= 2.0, 5= 1.0, 6= 0.5, 7= 0.2, 8= 0.1
0x03 | 0x001A | 0x0001 | 40026 | List | INPUTS – CONFIG. – INPUTS | Number of channels which will by used for measurement | 0=1. INP., 1=2. INP., 2=3. INP., 3=4. INP., 4=5. INP., 5=6. INP., 6=7. INP., 7=8. INP., 8=2. FAST, 9=3. FAST, 10=4. FAST
0x03 | 0x001C | 0x0001 | 40028 | List | INPUTS – CONFIG. – IN. MOD. | Overflow value | 0=CYKL., 1=SWITCH
0x03 | 0x001E | 0x0001 | 40030 | List | INPUTS – CONFIG. – SWITCH | Choice switch mode | 0=MANUAL, 1=AUTOM.
0x03 | 0x0020 | 0x0002 | 40032 | Float | INPUTS – CONFIG. – TIM. SW. | Switching time | 0,5..99,9
0x03 | 0x0022 | 0x0001 | 40034 | List | INPUTS – CONFIG. – INP. 1 – TYPE 1 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x0024 | 0x0001 | 40036 | List | INPUTS – CONFIG. – INP. 1 – MODE 1 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x0026 | 0x0001 | 40038 | List | INPUTS – CONFIG. – INP. 1 – CONE. 1 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0028 | 0x0002 | 40040 | Float | INPUTS – CONFIG. – INP. 1 – C.J. TEM. | CJC temperature | 0..99
0x03 | 0x002A | 0x0002 | 40042 | Float | INPUTS – CONFIG. – INP. 1 – AD. R. 1 | Offset value | 0..9999
0x03 | 0x002C | 0x0001 | 40044 | Char | INPUTS – CONFIG. – INP. 1 – LEAD. 1 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x002E | 0x0001 | 40046 | List | INPUTS – CONFIG. – INP. 2 – TYPE 2 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x0030 | 0x0001 | 40048 | List | INPUTS – CONFIG. – INP. 2 – MODE 2 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x0032 | 0x0001 | 40050 | List | INPUTS – CONFIG. – INP. 2 – CONE. 2 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0034 | 0x0002 | 40052 | Float | INPUTS – CONFIG. – INP. 2 – AD. R. 2 | Offset value | 0..9999
0x03 | 0x0036 | 0x0001 | 40054 | Char | INPUTS – CONFIG. – INP. 2 – LEAD. 2 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x0038 | 0x0001 | 40056 | List | INPUTS – CONFIG. – INP. 3 – TYPE 3 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x003A | 0x0001 | 40058 | List | INPUTS – CONFIG. – INP. 3 – MODE 3 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x003C | 0x0001 | 40060 | List | INPUTS – CONFIG. – INP. 3 – CONE. 3 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x003E | 0x0002 | 40062 | Float | INPUTS – CONFIG. – INP. 3 – AD. R. 3 | Offset value | 0..9999
0x03 | 0x0040 | 0x0001 | 40064 | Char | INPUTS – CONFIG. – INP. 3 – LEAD. 3 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x0042 | 0x0001 | 40066 | List | INPUTS – CONFIG. – INP. 4 – TYPE 4 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x0044 | 0x0001 | 40068 | List | INPUTS – CONFIG. – INP. 4 – MODE 4 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x0046 | 0x0001 | 40070 | List | INPUTS – CONFIG. – INP. 4 – CONE. 4 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0048 | 0x0002 | 40072 | Float | INPUTS – CONFIG. – INP. 4 – AD. R. 4 | Offset value | 0..9999
0x03 | 0x004A | 0x0001 | 40074 | Char | INPUTS – CONFIG. – INP. 4 – LEAD. 4 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x004C | 0x0001 | 40076 | List | INPUTS – CONFIG. – INP. 5 – TYPE 5 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x004E | 0x0001 | 40078 | List | INPUTS – CONFIG. – INP. 5 – MODE 5 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x0050 | 0x0001 | 40080 | List | INPUTS – CONFIG. – INP. 5 – CONE. 5 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0052 | 0x0002 | 40082 | Float | INPUTS – CONFIG. – INP. 5 – AD. R. 5 | Offset value | 0..9999
0x03 | 0x0054 | 0x0001 | 40084 | Char | INPUTS – CONFIG. – INP. 5 – LEAD. 5 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x0056 | 0x0001 | 40086 | List | INPUTS – CONFIG. – INP. 6 – TYPE 6 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x0058 | 0x0001 | 40088 | List | INPUTS – CONFIG. – INP. 6 – MODE 6 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x005A | 0x0001 | 40090 | List | INPUTS – CONFIG. – INP. 6 – CONE. 6 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x005C | 0x0002 | 40092 | Float | INPUTS – CONFIG. – INP. 6 – AD. R. 6 | Offset value | 0..9999
0x03 | 0x005E | 0x0001 | 40094 | Char | INPUTS – CONFIG. – INP. 6 – LEAD. 6 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x0060 | 0x0001 | 40096 | List | INPUTS – CONFIG. – INP. 7 – TYPE 7 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x0062 | 0x0001 | 40098 | List | INPUTS – CONFIG. – INP. 7 – MODE 7 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x0064 | 0x0001 | 40100 | List | INPUTS – CONFIG. – INP. 7 – CONE. 7 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0066 | 0x0002 | 40102 | Float | INPUTS – CONFIG. – INP. 7 – AD. R. 7 | Offset value | 0..9999
0x03 | 0x0068 | 0x0001 | 40104 | Char | INPUTS – CONFIG. – INP. 7 – LEAD. 7 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x006A | 0x0001 | 40106 | List | INPUTS – CONFIG. – INP. 8 – TYPE 8 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x03 | 0x006C | 0x0001 | 40108 | List | INPUTS – CONFIG. – INP. 8 – MODE 8 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x03 | 0x006E | 0x0001 | 40110 | List | INPUTS – CONFIG. – INP. 8 – CONE. 8 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x03 | 0x0070 | 0x0002 | 40112 | Float | INPUTS – CONFIG. – INP. 8 – AD. R. 8 | Offset value | 0..9999
0x03 | 0x0072 | 0x0001 | 40114 | Char | INPUTS – CONFIG. – INP. 8 – LEAD. 8 | 3L. GRID. Calculation of 3-phase power output - the value is calculated with the assumption of a balanced demand in all phases - for P, S Q the value is multiplied by 3 | ..
0x03 | 0x0078 | 0x0001 | 40120 | List | INPUTS – EXT. IN. – EXT. 1 | Setting auxiliary input 1 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x03 | 0x007A | 0x0001 | 40122 | List | INPUTS – EXT. IN. – EXT. 2 | Setting auxiliary input 2 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x03 | 0x007C | 0x0001 | 40124 | List | INPUTS – EXT. IN. – EXT. 3 | Setting auxiliary input 3 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x03 | 0x007E | 0x0001 | 40126 | List | INPUTS – EXT. IN. – M. HOLD | Range of instrument blocation | 0=DISPL., 1=DIS.+A.O., 2=D.+A.O.+L., 3=ALL
0x03 | 0x0080 | 0x0001 | 40128 | List | INPUTS – KEYS – LEFT – FN. LE. | Function setting for key LEFT | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x03 | 0x0082 | 0x0001 | 40130 | List | INPUTS – KEYS – LEFT – TMP. LE. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0084 | 0x0001 | 40132 | List | INPUTS – KEYS – LEFT – MNU. LE. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x03 | 0x0086 | 0x0001 | 40134 | List | INPUTS – KEYS – DOWN – FN. DO. | Function setting for key DOWN | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x03 | 0x0088 | 0x0001 | 40136 | List | INPUTS – KEYS – DOWN – TMP. DO. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x008A | 0x0001 | 40138 | List | INPUTS – KEYS – DOWN – MNU. DO. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x03 | 0x008C | 0x0001 | 40140 | List | INPUTS – KEYS – UP – FN. UP | Function setting for key UP | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x03 | 0x008E | 0x0001 | 40142 | List | INPUTS – KEYS – UP – TMP. UP | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0090 | 0x0001 | 40144 | List | INPUTS – KEYS – UP – MNU. UP | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x03 | 0x0092 | 0x0001 | 40146 | List | INPUTS – KEYS – ENTER – FN. EN. | Function setting for key ENTER | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x03 | 0x0094 | 0x0001 | 40148 | List | INPUTS – KEYS – ENTER – TMP. EN. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0096 | 0x0001 | 40150 | List | INPUTS – KEYS – ENTER – MNU. EN. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x03 | 0x0098 | 0x0002 | 40152 | Float | CHANNE. – CHAN. A – SET. A – MIN A | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x009A | 0x0002 | 40154 | Float | CHANNE. – CHAN. A – SET. A – MAX A | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x009C | 0x0002 | 40156 | Float | CHANNE. – CHAN. A – SET. A – P.TAR. A | Handy set value of tare | -999..9999
0x03 | 0x009E | 0x0001 | 40158 | List | CHANNE. – CHAN. A – FILTER – MOD. F.A | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x00A0 | 0x0002 | 40160 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 0,001..9999
0x03 | 0x00A2 | 0x0002 | 40162 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 2..Filter2Delka
0x03 | 0x00A4 | 0x0002 | 40164 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 2..Filter3Delka
0x03 | 0x00A6 | 0x0001 | 40166 | List | CHANNE. – CHAN. A – FORM. A | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x00A8 | 0x0001 | 40168 | Desc | CHANNE. – CHAN. A – DESC. A | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x00AA | 0x0001 | 40170 | List | CHANNE. – CHAN. A – SWIT. A | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x00AC | 0x0001 | 40172 | List | CHANNE. – CHAN. A – LOG. A – SAVE A | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x00AE | 0x0002 | 40174 | Float | CHANNE. – CHAN. A – LOG. A – FROM A | Lower limit of data storing range | -999..9999
0x03 | 0x00B0 | 0x0002 | 40176 | Float | CHANNE. – CHAN. A – LOG. A – TO A | Upper limit of data storing range | -999..9999
0x03 | 0x00B2 | 0x0002 | 40178 | Float | CHANNE. – CHAN. B – SET. B – MIN B | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x00B4 | 0x0002 | 40180 | Float | CHANNE. – CHAN. B – SET. B – MAX B | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x00B6 | 0x0002 | 40182 | Float | CHANNE. – CHAN. B – SET. B – P.TAR. B | Handy set value of tare | -999..9999
0x03 | 0x00B8 | 0x0001 | 40184 | List | CHANNE. – CHAN. B – FILTER – MOD. F.B | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x00BA | 0x0002 | 40186 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 0,001..9999
0x03 | 0x00BC | 0x0002 | 40188 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 2..Filter2Delka
0x03 | 0x00BE | 0x0002 | 40190 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 2..Filter3Delka
0x03 | 0x00C0 | 0x0001 | 40192 | List | CHANNE. – CHAN. B – FORM. B | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x00C2 | 0x0001 | 40194 | Desc | CHANNE. – CHAN. B – DESC. B | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x00C4 | 0x0001 | 40196 | List | CHANNE. – CHAN. B – SWIT. B | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x00C6 | 0x0001 | 40198 | List | CHANNE. – CHAN. B – LOG. B – SAVE B | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x00C8 | 0x0002 | 40200 | Float | CHANNE. – CHAN. B – LOG. B – FROM B | Lower limit of data storing range | -999..9999
0x03 | 0x00CA | 0x0002 | 40202 | Float | CHANNE. – CHAN. B – LOG. B – TO B | Upper limit of data storing range | -999..9999
0x03 | 0x00CC | 0x0002 | 40204 | Float | CHANNE. – CHAN. C – SET. C – MIN C | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x00CE | 0x0002 | 40206 | Float | CHANNE. – CHAN. C – SET. C – MAX C | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x00D0 | 0x0002 | 40208 | Float | CHANNE. – CHAN. C – SET. C – P.TAR. C | Handy set value of tare | -999..9999
0x03 | 0x00D2 | 0x0001 | 40210 | List | CHANNE. – CHAN. C – FILTER – MOD. F.C | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x00D4 | 0x0002 | 40212 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 0,001..9999
0x03 | 0x00D6 | 0x0002 | 40214 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 2..Filter2Delka
0x03 | 0x00D8 | 0x0002 | 40216 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 2..Filter3Delka
0x03 | 0x00DA | 0x0001 | 40218 | List | CHANNE. – CHAN. C – FORM. C | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x00DC | 0x0001 | 40220 | Desc | CHANNE. – CHAN. C – DESC. C | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x00DE | 0x0001 | 40222 | List | CHANNE. – CHAN. C – SWIT. C | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x00E0 | 0x0001 | 40224 | List | CHANNE. – CHAN. C – LOG. C – SAVE C | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x00E2 | 0x0002 | 40226 | Float | CHANNE. – CHAN. C – LOG. C – FROM C | Lower limit of data storing range | -999..9999
0x03 | 0x00E4 | 0x0002 | 40228 | Float | CHANNE. – CHAN. C – LOG. C – TO C | Upper limit of data storing range | -999..9999
0x03 | 0x00E6 | 0x0002 | 40230 | Float | CHANNE. – CHAN. D – SET. D – MIN D | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x00E8 | 0x0002 | 40232 | Float | CHANNE. – CHAN. D – SET. D – MAX D | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x00EA | 0x0002 | 40234 | Float | CHANNE. – CHAN. D – SET. D – P.TAR. D | Handy set value of tare | -999..9999
0x03 | 0x00EC | 0x0001 | 40236 | List | CHANNE. – CHAN. D – FILTER – MOD. F.D | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x00EE | 0x0002 | 40238 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 0,001..9999
0x03 | 0x00F0 | 0x0002 | 40240 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 2..Filter2Delka
0x03 | 0x00F2 | 0x0002 | 40242 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 2..Filter3Delka
0x03 | 0x00F4 | 0x0001 | 40244 | List | CHANNE. – CHAN. D – FORM. D | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x00F6 | 0x0001 | 40246 | Desc | CHANNE. – CHAN. D – DESC. D | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x00F8 | 0x0001 | 40248 | List | CHANNE. – CHAN. D – SWIT. D | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x00FA | 0x0001 | 40250 | List | CHANNE. – CHAN. D – LOG. D – SAVE D | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x00FC | 0x0002 | 40252 | Float | CHANNE. – CHAN. D – LOG. D – FROM D | Lower limit of data storing range | -999..9999
0x03 | 0x00FE | 0x0002 | 40254 | Float | CHANNE. – CHAN. D – LOG. D – TO D | Upper limit of data storing range | -999..9999
0x03 | 0x0100 | 0x0002 | 40256 | Float | CHANNE. – CHAN. E – SET. E – MIN E | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x0102 | 0x0002 | 40258 | Float | CHANNE. – CHAN. E – SET. E – MAX E | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x0104 | 0x0002 | 40260 | Float | CHANNE. – CHAN. E – SET. E – P.TAR. E | Handy set value of tare | -999..9999
0x03 | 0x0106 | 0x0001 | 40262 | List | CHANNE. – CHAN. E – FILTER – MOD. F.E | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x0108 | 0x0002 | 40264 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 0,001..9999
0x03 | 0x010A | 0x0002 | 40266 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 2..Filter2Delka
0x03 | 0x010C | 0x0002 | 40268 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 2..Filter3Delka
0x03 | 0x010E | 0x0001 | 40270 | List | CHANNE. – CHAN. E – FORM. E | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x0110 | 0x0001 | 40272 | Desc | CHANNE. – CHAN. E – DESC. E | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x0112 | 0x0001 | 40274 | List | CHANNE. – CHAN. E – SWIT. E | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x0114 | 0x0001 | 40276 | List | CHANNE. – CHAN. E – LOG. E – SAVE E | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x0116 | 0x0002 | 40278 | Float | CHANNE. – CHAN. E – LOG. E – FROM E | Lower limit of data storing range | -999..9999
0x03 | 0x0118 | 0x0002 | 40280 | Float | CHANNE. – CHAN. E – LOG. E – TO E | Upper limit of data storing range | -999..9999
0x03 | 0x011A | 0x0002 | 40282 | Float | CHANNE. – CHAN. F – SET. F – MIN F | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x011C | 0x0002 | 40284 | Float | CHANNE. – CHAN. F – SET. F – MAX F | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x011E | 0x0002 | 40286 | Float | CHANNE. – CHAN. F – SET. F – P.TAR. F | Handy set value of tare | -999..9999
0x03 | 0x0120 | 0x0001 | 40288 | List | CHANNE. – CHAN. F – FILTER – MOD. F.F | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x0122 | 0x0002 | 40290 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 0,001..9999
0x03 | 0x0124 | 0x0002 | 40292 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 2..Filter2Delka
0x03 | 0x0126 | 0x0002 | 40294 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 2..Filter3Delka
0x03 | 0x0128 | 0x0001 | 40296 | List | CHANNE. – CHAN. F – FORM. F | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x012A | 0x0001 | 40298 | Desc | CHANNE. – CHAN. F – DESC. F | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x012C | 0x0001 | 40300 | List | CHANNE. – CHAN. F – SWIT. F | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x012E | 0x0001 | 40302 | List | CHANNE. – CHAN. F – LOG. F – SAVE F | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x0130 | 0x0002 | 40304 | Float | CHANNE. – CHAN. F – LOG. F – FROM F | Lower limit of data storing range | -999..9999
0x03 | 0x0132 | 0x0002 | 40306 | Float | CHANNE. – CHAN. F – LOG. F – TO F | Upper limit of data storing range | -999..9999
0x03 | 0x0134 | 0x0002 | 40308 | Float | CHANNE. – CHAN. G – SET. G – MIN G | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x0136 | 0x0002 | 40310 | Float | CHANNE. – CHAN. G – SET. G – MAX G | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x0138 | 0x0002 | 40312 | Float | CHANNE. – CHAN. G – SET. G – P.TAR. G | Handy set value of tare | -999..9999
0x03 | 0x013A | 0x0001 | 40314 | List | CHANNE. – CHAN. G – FILTER – MOD. F.G | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x013C | 0x0002 | 40316 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 0,001..9999
0x03 | 0x013E | 0x0002 | 40318 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 2..Filter2Delka
0x03 | 0x0140 | 0x0002 | 40320 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 2..Filter3Delka
0x03 | 0x0142 | 0x0001 | 40322 | List | CHANNE. – CHAN. G – FORM. G | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x0144 | 0x0001 | 40324 | Desc | CHANNE. – CHAN. G – DESC. G | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x0146 | 0x0001 | 40326 | List | CHANNE. – CHAN. G – SWIT. G | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x0148 | 0x0001 | 40328 | List | CHANNE. – CHAN. G – LOG. G – SAVE G | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x014A | 0x0002 | 40330 | Float | CHANNE. – CHAN. G – LOG. G – FROM G | Lower limit of data storing range | -999..9999
0x03 | 0x014C | 0x0002 | 40332 | Float | CHANNE. – CHAN. G – LOG. G – TO G | Upper limit of data storing range | -999..9999
0x03 | 0x014E | 0x0002 | 40334 | Float | CHANNE. – CHAN. H – SET. H – MIN H | Value corresponding w/ minimum value of input range | -999..9999
0x03 | 0x0150 | 0x0002 | 40336 | Float | CHANNE. – CHAN. H – SET. H – MAX H | Value corresponding w/ maximum value of input range | -999..9999
0x03 | 0x0152 | 0x0002 | 40338 | Float | CHANNE. – CHAN. H – SET. H – P.TAR. H | Handy set value of tare | -999..9999
0x03 | 0x0154 | 0x0001 | 40340 | List | CHANNE. – CHAN. H – FILTER – MOD. F.H | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x03 | 0x0156 | 0x0002 | 40342 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 0,001..9999
0x03 | 0x0158 | 0x0002 | 40344 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 2..Filter2Delka
0x03 | 0x015A | 0x0002 | 40346 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 2..Filter3Delka
0x03 | 0x015C | 0x0001 | 40348 | List | CHANNE. – CHAN. H – FORM. H | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x015E | 0x0001 | 40350 | Desc | CHANNE. – CHAN. H – DESC. H | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x0160 | 0x0001 | 40352 | List | CHANNE. – CHAN. H – SWIT. H | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x03 | 0x0162 | 0x0001 | 40354 | List | CHANNE. – CHAN. H – LOG. H – SAVE H | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x03 | 0x0164 | 0x0002 | 40356 | Float | CHANNE. – CHAN. H – LOG. H – FROM H | Lower limit of data storing range | -999..9999
0x03 | 0x0166 | 0x0002 | 40358 | Float | CHANNE. – CHAN. H – LOG. H – TO H | Upper limit of data storing range | -999..9999
0x03 | 0x0168 | 0x0001 | 40360 | List | CHANNE. – MAT. FN. – INP. M. | Input value | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x016A | 0x0001 | 40362 | List | CHANNE. – MAT. FN. – MATH. F. | Mathematic function select | 0=MULTIN., 1=1/MUL., 2=LOGAR., 3=EXPON., 4=POWER, 5=ROOT, 6=SIN X, 7=SUMA., 8=DIVIDE., 9=QUADR. .
0x03 | 0x016C | 0x0002 | 40364 | Float | CHANNE. – MAT. FN. – CON. A | Constant A | -999..9999
0x03 | 0x016E | 0x0002 | 40366 | Float | CHANNE. – MAT. FN. – CON. B | Constant B | -999..9999
0x03 | 0x0170 | 0x0002 | 40368 | Float | CHANNE. – MAT. FN. – CON. C | Constant C | -999..9999
0x03 | 0x0172 | 0x0002 | 40370 | Float | CHANNE. – MAT. FN. – CON. D | Constant D | -999..9999
0x03 | 0x0174 | 0x0002 | 40372 | Float | CHANNE. – MAT. FN. – CON. E | Constant E | -999..9999
0x03 | 0x0176 | 0x0002 | 40374 | Float | CHANNE. – MAT. FN. – CON. F | Constant F | -999..9999
0x03 | 0x0178 | 0x0002 | 40376 | Float | CHANNE. – MAT. FN. – CON. G | Constant G | -999..9999
0x03 | 0x017A | 0x0002 | 40378 | Float | CHANNE. – MAT. FN. – CON. H | Constant H | -999..9999
0x03 | 0x017C | 0x0002 | 40380 | Float | CHANNE. – MAT. FN. – CON. I | Constant I | -999..9999
0x03 | 0x017E | 0x0002 | 40382 | Float | CHANNE. – MAT. FN. – CON. J | Constant J | -999..9999
0x03 | 0x0180 | 0x0001 | 40384 | List | CHANNE. – MAT. FN. – FORM. M | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x03 | 0x0182 | 0x0001 | 40386 | Desc | CHANNE. – MAT. FN. – DESC. M | Description setting (measured quantity unit) | Char(32)..Char(127)
0x03 | 0x0184 | 0x0001 | 40388 | List | CHANNE. – MAT. FN. – SWIT. M | Setting of display cycle switching | 0=NO, 1=YES
0x03 | 0x018C | 0x0001 | 40396 | List | CHANNE. – MIN.MAX – INP. M.M. | Input value for minimum and maximum | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x018E | 0x0001 | 40398 | CheckBox | CHANNE. – MIN.MAX – CH. M.M. |  | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x0190 | 0x0001 | 40400 | List | OUTPUT. – MEMORY – REWRIT. | Oldest values are rewrited with newest | 0=NO, 1=YES
0x03 | 0x0192 | 0x0002 | 40402 | Long | OUTPUT. – MEMORY – START | Start of data storing interval | 0..1522491
0x03 | 0x0194 | 0x0002 | 40404 | Long | OUTPUT. – MEMORY – STOP | End of datastoring interval | 0..1522491
0x03 | 0x0196 | 0x0002 | 40406 | Long | OUTPUT. – MEMORY – PERIOD. | Period of datastoring | 0..1522491
0x03 | 0x0198 | 0x0002 | 40408 | Float | OUTPUT. – MEMORY – TRIGG. | Size of pretrigger area of memory | 0..100
0x03 | 0x019A | 0x0001 | 40410 | List | OUTPUT. – LIMITS – LIM 1 – INP. L.1 | Input value for limit 1 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x019C | 0x0001 | 40412 | CheckBox | OUTPUT. – LIMITS – LIM 1 – CH. L.1 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x019E | 0x0001 | 40414 | List | OUTPUT. – LIMITS – LIM 1 – MOD. L.1 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x01A0 | 0x0001 | 40416 | List | OUTPUT. – LIMITS – LIM 1 – TYP. L.1 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x01A2 | 0x0002 | 40418 | Float | OUTPUT. – LIMITS – LIM 1 – LIM. L.1 | Limit value | -999..9999
0x03 | 0x01A4 | 0x0002 | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – HYS. L.1 | Hysteresis | 0..9999
0x03 | *0x01A2 | 0x0002 | 40418 | Float | OUTPUT. – LIMITS – LIM 1 – ON L.1 | Limit value | -999..9999
0x03 | *0x01A4 | 0x0002 | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – OFF L.1 | Hysteresis | -999..9999
0x03 | *0x01A4 | 0x0002 | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – PER. L.1 | Hysteresis | 0,001..9999
0x03 | 0x01AC | 0x0002 | 40428 | Float | OUTPUT. – LIMITS – LIM 1 – TIM. L.1 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x01AE | 0x0001 | 40430 | List | OUTPUT. – LIMITS – LIM 2 – INP. L.2 | Input value for limit 2 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x01B0 | 0x0001 | 40432 | CheckBox | OUTPUT. – LIMITS – LIM 2 – CH. L.2 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x01B2 | 0x0001 | 40434 | List | OUTPUT. – LIMITS – LIM 2 – MOD. L.2 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x01B4 | 0x0001 | 40436 | List | OUTPUT. – LIMITS – LIM 2 – TYP. L.2 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x01B6 | 0x0002 | 40438 | Float | OUTPUT. – LIMITS – LIM 2 – LIM. L.2 | Limit value | -999..9999
0x03 | 0x01B8 | 0x0002 | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – HYS. L.2 | Hysteresis | 0..9999
0x03 | *0x01B6 | 0x0002 | 40438 | Float | OUTPUT. – LIMITS – LIM 2 – ON L.2 | Limit value | -999..9999
0x03 | *0x01B8 | 0x0002 | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – OFF L.2 | Hysteresis | -999..9999
0x03 | *0x01B8 | 0x0002 | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – PER. L.2 | Hysteresis | 0,001..9999
0x03 | 0x01C0 | 0x0002 | 40448 | Float | OUTPUT. – LIMITS – LIM 2 – TIM. L.2 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x01C2 | 0x0001 | 40450 | List | OUTPUT. – LIMITS – LIM 3 – INP. L.3 | Input value for limit 3 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x01C4 | 0x0001 | 40452 | CheckBox | OUTPUT. – LIMITS – LIM 3 – CH. L.3 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x01C6 | 0x0001 | 40454 | List | OUTPUT. – LIMITS – LIM 3 – MOD. L.3 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x01C8 | 0x0001 | 40456 | List | OUTPUT. – LIMITS – LIM 3 – TYP. L.3 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x01CA | 0x0002 | 40458 | Float | OUTPUT. – LIMITS – LIM 3 – LIM. L.3 | Limit value | -999..9999
0x03 | 0x01CC | 0x0002 | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – HYS. L.3 | Hysteresis | 0..9999
0x03 | *0x01CA | 0x0002 | 40458 | Float | OUTPUT. – LIMITS – LIM 3 – ON L.3 | Limit value | -999..9999
0x03 | *0x01CC | 0x0002 | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – OFF L.3 | Hysteresis | -999..9999
0x03 | *0x01CC | 0x0002 | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – PER. L.3 | Hysteresis | 0,001..9999
0x03 | 0x01D4 | 0x0002 | 40468 | Float | OUTPUT. – LIMITS – LIM 3 – TIM. L.3 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x01D6 | 0x0001 | 40470 | List | OUTPUT. – LIMITS – LIM 4 – INP. L.4 | Input value for limit 4 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x01D8 | 0x0001 | 40472 | CheckBox | OUTPUT. – LIMITS – LIM 4 – CH. L.4 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x01DA | 0x0001 | 40474 | List | OUTPUT. – LIMITS – LIM 4 – MOD. L.4 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x01DC | 0x0001 | 40476 | List | OUTPUT. – LIMITS – LIM 4 – TYP. L.4 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x01DE | 0x0002 | 40478 | Float | OUTPUT. – LIMITS – LIM 4 – LIM. L.4 | Limit value | -999..9999
0x03 | 0x01E0 | 0x0002 | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – HYS. L.4 | Hysteresis | 0..9999
0x03 | *0x01DE | 0x0002 | 40478 | Float | OUTPUT. – LIMITS – LIM 4 – ON L.4 | Limit value | -999..9999
0x03 | *0x01E0 | 0x0002 | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – OFF L.4 | Hysteresis | -999..9999
0x03 | *0x01E0 | 0x0002 | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – PER. L.4 | Hysteresis | 0,001..9999
0x03 | 0x01E8 | 0x0002 | 40488 | Float | OUTPUT. – LIMITS – LIM 4 – TIM. L.4 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x01EA | 0x0001 | 40490 | List | OUTPUT. – LIMITS – LIM 5 – INP. L.5 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x01EC | 0x0001 | 40492 | CheckBox | OUTPUT. – LIMITS – LIM 5 – CH. L.5 |  | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x01EE | 0x0001 | 40494 | List | OUTPUT. – LIMITS – LIM 5 – MOD. L.5 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x01F0 | 0x0001 | 40496 | List | OUTPUT. – LIMITS – LIM 5 – TYP. L.5 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x01F2 | 0x0002 | 40498 | Float | OUTPUT. – LIMITS – LIM 5 – LIM. L.5 | Limit value | -999..9999
0x03 | 0x01F4 | 0x0002 | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – HYS. L.5 | Hysteresis | 0..9999
0x03 | *0x01F2 | 0x0002 | 40498 | Float | OUTPUT. – LIMITS – LIM 5 – ON L.5 | Limit value | -999..9999
0x03 | *0x01F4 | 0x0002 | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – OFF L.5 | Hysteresis | -999..9999
0x03 | *0x01F4 | 0x0002 | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – PER. L.5 | Hysteresis | 0,001..9999
0x03 | 0x01FC | 0x0002 | 40508 | Float | OUTPUT. – LIMITS – LIM 5 – TIM. L.5. | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x01FE | 0x0001 | 40510 | List | OUTPUT. – LIMITS – LIM 6 – INP. L.6 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0200 | 0x0001 | 40512 | CheckBox | OUTPUT. – LIMITS – LIM 6 – CH. L.6 |  | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x0202 | 0x0001 | 40514 | List | OUTPUT. – LIMITS – LIM 6 – MOD. L.6 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x0204 | 0x0001 | 40516 | List | OUTPUT. – LIMITS – LIM 6 – TYP. L.6 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x0206 | 0x0002 | 40518 | Float | OUTPUT. – LIMITS – LIM 6 – LIM. L.6 | Limit value | -999..9999
0x03 | 0x0208 | 0x0002 | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – HYS. L.6 | Hysteresis | 0..9999
0x03 | *0x0206 | 0x0002 | 40518 | Float | OUTPUT. – LIMITS – LIM 6 – ON L.6 | Limit value | -999..9999
0x03 | *0x0208 | 0x0002 | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – OFF L.6 | Hysteresis | -999..9999
0x03 | *0x0208 | 0x0002 | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – PER. L.6 | Hysteresis | 0,001..9999
0x03 | 0x0210 | 0x0002 | 40528 | Float | OUTPUT. – LIMITS – LIM 6 – TIM. L.6 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x0212 | 0x0001 | 40530 | List | OUTPUT. – LIMITS – LIM 7 – INP. L.7 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0214 | 0x0001 | 40532 | CheckBox | OUTPUT. – LIMITS – LIM 7 – CH. L.7 |  | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x0216 | 0x0001 | 40534 | List | OUTPUT. – LIMITS – LIM 7 – MOD. L.7 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x0218 | 0x0001 | 40536 | List | OUTPUT. – LIMITS – LIM 7 – TYP. L.7 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x021A | 0x0002 | 40538 | Float | OUTPUT. – LIMITS – LIM 7 – LIM. L.7 | Limit value | -999..9999
0x03 | 0x021C | 0x0002 | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – HYS. L.7 | Hysteresis | 0..9999
0x03 | *0x021A | 0x0002 | 40538 | Float | OUTPUT. – LIMITS – LIM 7 – ON L.7 | Limit value | -999..9999
0x03 | *0x021C | 0x0002 | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – OFF L.7 | Hysteresis | -999..9999
0x03 | *0x021C | 0x0002 | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – PER. L.7 | Hysteresis | 0,001..9999
0x03 | 0x0224 | 0x0002 | 40548 | Float | OUTPUT. – LIMITS – LIM 7 – TIM. L.7 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x0226 | 0x0001 | 40550 | List | OUTPUT. – LIMITS – LIM 8 – INP. L.8 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0228 | 0x0001 | 40552 | CheckBox | OUTPUT. – LIMITS – LIM 8 – CH. L.8 |  | 0..15(4 channels), 0..255(8 channels)
0x03 | 0x022A | 0x0001 | 40554 | List | OUTPUT. – LIMITS – LIM 8 – MOD. L.8 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x03 | 0x022C | 0x0001 | 40556 | List | OUTPUT. – LIMITS – LIM 8 – TYP. L.8 | Limit type | 0=CLOSE, 1=OPEN
0x03 | 0x022E | 0x0002 | 40558 | Float | OUTPUT. – LIMITS – LIM 8 – LIM. L.8 | Limit value | -999..9999
0x03 | 0x0230 | 0x0002 | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – HYS. L.8 | Hysteresis | 0..9999
0x03 | *0x022E | 0x0002 | 40558 | Float | OUTPUT. – LIMITS – LIM 8 – ON L.8 | Limit value | -999..9999
0x03 | *0x0230 | 0x0002 | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – OFF L.8 | Hysteresis | -999..9999
0x03 | *0x0230 | 0x0002 | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – PER. L.8 | Hysteresis | 0,001..9999
0x03 | 0x0238 | 0x0002 | 40568 | Float | OUTPUT. – LIMITS – LIM 8 – TIM. L.8 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x03 | 0x023A | 0x0001 | 40570 | List | OUTPUT. – DATA – BAUD | Transfer rate | 0= 600, 1= 1200, 2= 2400, 3= 4800, 4= 9600, 5= 19200, 6= 38400, 7= 57600, 8=115200, 9=230400
0x03 | 0x023C | 0x0002 | 40572 | Float | OUTPUT. – DATA – ADDR. | Instrument address | 0..31
0x03 | 0x0240 | 0x0002 | 40576 | Float | OUTPUT. – DATA – ADR. P.B. | Returns PROFIBUS address and enable its editation in menu. | 0..125
0x03 | 0x0242 | 0x0001 | 40578 | List | OUTPUT. – DATA – PROT. | Communication protocol | 0=ASCII, 1=M. BUS, 2=MODBUS
0x03 | 0x0244 | 0x0001 | 40580 | List | OUTPUT. – AN. OUT. – INP. A.O. | Input value for analog output | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x0246 | 0x0001 | 40582 | List | OUTPUT. – AN. OUT. – TYP. A.O. | Type | 0=0-20mA, 1=Er4- T, 2=4-20 T, 3=Er4-20, 4=4-20mA, 5=0- 5mA, 6=0- 2 V, 7=0- 5 V, 8=0-10 V, 9=+-10 V
0x03 | 0x0248 | 0x0002 | 40584 | Float | OUTPUT. – AN. OUT. – MIN A.O. | Minimum display value corresponding with bottom value of selected A.O. range | -999..9999
0x03 | 0x024A | 0x0002 | 40586 | Float | OUTPUT. – AN. OUT. – MAX A.O. | Maximum display value corresponding with top value of selected A.O. range | -999..9999
0x03 | 0x024C | 0x0001 | 40588 | List | OUTPUT. – DISP. – PERM. | Value to be permanent displayed | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x03 | 0x024E | 0x0001 | 40590 | List | OUTPUT. – DISP. – CHAN. | This item intend actual channel number showing. | 0=NO, 1=NUMBER., 2=CHARS.
0x03 | 0x0250 | 0x0001 | 40592 | List | OUTPUT. – DISP. – BRIGHT | Brightness of displey | 0= 0%, 1= 25%, 2= 50%, 3= 75%, 4= 100%
0x03 | 0x0260 | 0x0001 | 40608 | List | SERVIC. – LANG. | Language variation of menu | 0=CZECH, 1=ENGL.
Nastavení hodnot menu / Setting menu item values
0x06 | 0x0018 | 0x00XX | 40024 | List | INPUTS – CONFIG. – READ./S | Measuring rate | 0= 40.0, 1= 20.0, 2= 10.0, 3= 5.0, 4= 2.0, 5= 1.0, 6= 0.5, 7= 0.2, 8= 0.1
0x06 | 0x001A | 0x00XX | 40026 | List | INPUTS – CONFIG. – INPUTS | Number of channels which will by used for measurement | 0=1. INP., 1=2. INP., 2=3. INP., 3=4. INP., 4=5. INP., 5=6. INP., 6=7. INP., 7=8. INP., 8=2. FAST, 9=3. FAST, 10=4. FAST
0x06 | 0x001C | 0x00XX | 40028 | List | INPUTS – CONFIG. – IN. MOD. | Overflow value | 0=CYKL., 1=SWITCH
0x06 | 0x001E | 0x00XX | 40030 | List | INPUTS – CONFIG. – SWITCH | Choice switch mode | 0=MANUAL, 1=AUTOM.
0x10 | 0x0020 | 0x000204EEMMMMMM | 40032 | Float | INPUTS – CONFIG. – TIM. SW. | Switching time | 0,5..99,9
0x06 | 0x0022 | 0x00XX | 40034 | List | INPUTS – CONFIG. – INP. 1 – TYPE 1 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x0024 | 0x00XX | 40036 | List | INPUTS – CONFIG. – INP. 1 – MODE 1 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x0026 | 0x00XX | 40038 | List | INPUTS – CONFIG. – INP. 1 – CONE. 1 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0028 | 0x000204EEMMMMMM | 40040 | Float | INPUTS – CONFIG. – INP. 1 – C.J. TEM. | CJC temperature | 0..99
0x10 | 0x002A | 0x000204EEMMMMMM | 40042 | Float | INPUTS – CONFIG. – INP. 1 – AD. R. 1 | Additional resistance in probe | 0..9999
0x06 | 0x002E | 0x00XX | 40046 | List | INPUTS – CONFIG. – INP. 2 – TYPE 2 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x0030 | 0x00XX | 40048 | List | INPUTS – CONFIG. – INP. 2 – MODE 2 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x0032 | 0x00XX | 40050 | List | INPUTS – CONFIG. – INP. 2 – CONE. 2 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0034 | 0x000204EEMMMMMM | 40052 | Float | INPUTS – CONFIG. – INP. 2 – AD. R. 2 | Additional resistance in probe | 0..9999
0x06 | 0x0038 | 0x00XX | 40056 | List | INPUTS – CONFIG. – INP. 3 – TYPE 3 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x003A | 0x00XX | 40058 | List | INPUTS – CONFIG. – INP. 3 – MODE 3 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x003C | 0x00XX | 40060 | List | INPUTS – CONFIG. – INP. 3 – CONE. 3 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x003E | 0x000204EEMMMMMM | 40062 | Float | INPUTS – CONFIG. – INP. 3 – AD. R. 3 | Additional resistance in probe | 0..9999
0x06 | 0x0042 | 0x00XX | 40066 | List | INPUTS – CONFIG. – INP. 4 – TYPE 4 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x0044 | 0x00XX | 40068 | List | INPUTS – CONFIG. – INP. 4 – MODE 4 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x0046 | 0x00XX | 40070 | List | INPUTS – CONFIG. – INP. 4 – CONE. 4 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0048 | 0x000204EEMMMMMM | 40072 | Float | INPUTS – CONFIG. – INP. 4 – AD. R. 4 | Additional resistance in probe | 0..9999
0x06 | 0x004C | 0x00XX | 40076 | List | INPUTS – CONFIG. – INP. 5 – TYPE 5 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x004E | 0x00XX | 40078 | List | INPUTS – CONFIG. – INP. 5 – MODE 5 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x0050 | 0x00XX | 40080 | List | INPUTS – CONFIG. – INP. 5 – CONE. 5 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0052 | 0x000204EEMMMMMM | 40082 | Float | INPUTS – CONFIG. – INP. 5 – AD. R. 5 | Additional resistance in probe | 0..9999
0x06 | 0x0056 | 0x00XX | 40086 | List | INPUTS – CONFIG. – INP. 6 – TYPE 6 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x0058 | 0x00XX | 40088 | List | INPUTS – CONFIG. – INP. 6 – MODE 6 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x005A | 0x00XX | 40090 | List | INPUTS – CONFIG. – INP. 6 – CONE. 6 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x005C | 0x000204EEMMMMMM | 40092 | Float | INPUTS – CONFIG. – INP. 6 – AD. R. 6 | Additional resistance in probe | 0..9999
0x06 | 0x0060 | 0x00XX | 40096 | List | INPUTS – CONFIG. – INP. 7 – TYPE 7 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x0062 | 0x00XX | 40098 | List | INPUTS – CONFIG. – INP. 7 – MODE 7 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x0064 | 0x00XX | 40100 | List | INPUTS – CONFIG. – INP. 7 – CONE. 7 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0066 | 0x000204EEMMMMMM | 40102 | Float | INPUTS – CONFIG. – INP. 7 – AD. R. 7 | Additional resistance in probe | 0..9999
0x06 | 0x006A | 0x00XX | 40106 | List | INPUTS – CONFIG. – INP. 8 – TYPE 8 | Instrument type | 0=DC, 1=PM, 2=OHM, 3=RTD-Pt, 4=RTD-Ni, 5=TC, 6=DU, 7=RTD-Cu
0x06 | 0x006C | 0x00XX | 40108 | List | INPUTS – CONFIG. – INP. 8 – MODE 8 | Instrument mode | 0= 60 mV, 1=150 mV, 2=300 mV, 3=1200mV, 4=0- 5mA, 5=0-20mA, 6=4-20mA, 7=0- 2 V, 8=0- 5 V, 9=0-10V, 10=0-40V, 11= 100 R, 12= 1 k, 13= 10 k, 14= 100 k, 15=EU-100, 16=EU-500, 17=EU-1K0, 18=US-100, 19=5.0- 1K, 20=6.2- 1k, 21=5.0-10K, 22=6.2-10k, 23=T/C B, 24=T/C E, 25=T/C J, 26=T/C K, 27=T/C N, 28=T/C R, 29=T/C S, 30=T/C T, 31=LIN.POT., 32=RU- 50, 33=RU-100, 34=428-50, 35=428-0.1, 36=426-50, 37=426-0.1, 38=T/C L, 39=Er4-20
0x06 | 0x006E | 0x00XX | 40110 | List | INPUTS – CONFIG. – INP. 8 – CONE. 8 | Sensor connection | 0=INT.1TC, 1=INT.2TC, 2=EXT.1TC, 3=EXT.2TC, 4=2-WIRE, 5=3-WIRE, 6=4-WIRE
0x10 | 0x0070 | 0x000204EEMMMMMM | 40112 | Float | INPUTS – CONFIG. – INP. 8 – AD. R. 8 | Additional resistance in probe | 0..9999
0x06 | 0x0078 | 0x00XX | 40120 | List | INPUTS – EXT. IN. – EXT. 1 | Setting auxiliary input 1 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x06 | 0x007A | 0x00XX | 40122 | List | INPUTS – EXT. IN. – EXT. 2 | Setting auxiliary input 2 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x06 | 0x007C | 0x00XX | 40124 | List | INPUTS – EXT. IN. – EXT. 3 | Setting auxiliary input 3 | 0=OFF, 1=HOLD, 2=LOCK. K., 3=B. PASS, 4=TARE A, 5=TARE B, 6=TARE C, 7=TARE D, 8=TARE E, 9=TARE F, 10=TARE G, 11=TARE H, 12=TAR.ALL, 13=TAR.ACT, 14=CL. M.M., 15=CLR. T.A, 16=CLR. T.B, 17=CLR. T.C, 18=CLR. T.D, 19=CLR. T.E, 20=CLR. T.F, 21=CLR. T.G, 22=CLR. T.H, 23=C.T. ALL, 24=C.T. ACT, 25=SWCH. 1, 26=SWCH. 2, 27=SWCH. 3, 28=SAVE, 29=CL. MEM.
0x06 | 0x007E | 0x00XX | 40126 | List | INPUTS – EXT. IN. – M. HOLD | Range of instrument blocation | 0=DISPL., 1=DIS.+A.O., 2=D.+A.O.+L., 3=ALL
0x06 | 0x0080 | 0x00XX | 40128 | List | INPUTS – KEYS – LEFT – FN. LE. | Function setting for key LEFT | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x06 | 0x0082 | 0x00XX | 40130 | List | INPUTS – KEYS – LEFT – TMP. LE. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0084 | 0x00XX | 40132 | List | INPUTS – KEYS – LEFT – MNU. LE. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x06 | 0x0086 | 0x00XX | 40134 | List | INPUTS – KEYS – DOWN – FN. DO. | Function setting for key DOWN | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x06 | 0x0088 | 0x00XX | 40136 | List | INPUTS – KEYS – DOWN – TMP. DO. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x008A | 0x00XX | 40138 | List | INPUTS – KEYS – DOWN – MNU. DO. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x06 | 0x008C | 0x00XX | 40140 | List | INPUTS – KEYS – UP – FN. UP | Function setting for key UP | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x06 | 0x008E | 0x00XX | 40142 | List | INPUTS – KEYS – UP – TMP. UP | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0090 | 0x00XX | 40144 | List | INPUTS – KEYS – UP – MNU. UP | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x06 | 0x0092 | 0x00XX | 40146 | List | INPUTS – KEYS – ENTER – FN. EN. | Function setting for key ENTER | 0=NO, 1=CL. M.M., 2=CLR. T.A, 3=CLR. T.C, 4=CLR. T.C, 5=CLR. T.D, 6=CLR. T.E, 7=CLR. T.F, 8=CLR. T.G, 9=CLR. T.H, 10=C.T. ALL, 11=C.T. ACT, 12=MENU, 13=TEMP. V., 14=TARE A, 15=TARE B, 16=TARE C, 17=TARE D, 18=TARE E, 19=TARE F, 20=TARE G, 21=TARE H, 22=TAR.ALL, 23=TAR.ACT, 24=SWCH. 1, 25=SWCH.-1, 26=SAVE, 27=CL. MEM.
0x06 | 0x0094 | 0x00XX | 40148 | List | INPUTS – KEYS – ENTER – TMP. EN. | Choice of value for temporary show | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0096 | 0x00XX | 40150 | List | INPUTS – KEYS – ENTER – MNU. EN. | Choice of menu item | 0=LIM. 1, 1=LIM. 2, 2=LIM. 3, 3=LIM. 4, 4=LIM. 5, 5=LIM. 6, 6=LIM. 7, 7=LIM. 8
0x10 | 0x0098 | 0x000204EEMMMMMM | 40152 | Float | CHANNE. – CHAN. A – SET. A – MIN A | Scale | -999..9999
0x10 | 0x009A | 0x000204EEMMMMMM | 40154 | Float | CHANNE. – CHAN. A – SET. A – MAX A | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x009C | 0x000204EEMMMMMM | 40156 | Float | CHANNE. – CHAN. A – SET. A – P.TAR. A | Handy set value of tare | -999..9999
0x06 | 0x009E | 0x00XX | 40158 | List | CHANNE. – CHAN. A – FILTER – MOD. F.A | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x00A0 | 0x000204EEMMMMMM | 40160 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 0,001..9999
0x10 | 0x00A2 | 0x000204EEMMMMMM | 40162 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 2..Filter2Delka
0x10 | 0x00A4 | 0x000204EEMMMMMM | 40164 | Float | CHANNE. – CHAN. A – FILTER – CON. F.A | Filter constant | 2..Filter3Delka
0x06 | 0x00A6 | 0x00XX | 40166 | List | CHANNE. – CHAN. A – FORM. A | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x00A8 | 0xXXXX | 40168 | Desc | CHANNE. – CHAN. A – DESC. A | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x00AA | 0x00XX | 40170 | List | CHANNE. – CHAN. A – SWIT. A | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x00AC | 0x00XX | 40172 | List | CHANNE. – CHAN. A – LOG. A – SAVE A | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x00AE | 0x000204EEMMMMMM | 40174 | Float | CHANNE. – CHAN. A – LOG. A – FROM A | Lower limit of data storing range | -999..9999
0x10 | 0x00B0 | 0x000204EEMMMMMM | 40176 | Float | CHANNE. – CHAN. A – LOG. A – TO A | Upper limit of data storing range | -999..9999
0x10 | 0x00B2 | 0x000204EEMMMMMM | 40178 | Float | CHANNE. – CHAN. B – SET. B – MIN B | Scale | -999..9999
0x10 | 0x00B4 | 0x000204EEMMMMMM | 40180 | Float | CHANNE. – CHAN. B – SET. B – MAX B | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x00B6 | 0x000204EEMMMMMM | 40182 | Float | CHANNE. – CHAN. B – SET. B – P.TAR. B | Handy set value of tare | -999..9999
0x06 | 0x00B8 | 0x00XX | 40184 | List | CHANNE. – CHAN. B – FILTER – MOD. F.B | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x00BA | 0x000204EEMMMMMM | 40186 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 0,001..9999
0x10 | 0x00BC | 0x000204EEMMMMMM | 40188 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 2..Filter2Delka
0x10 | 0x00BE | 0x000204EEMMMMMM | 40190 | Float | CHANNE. – CHAN. B – FILTER – CON. F.B | Filter constant | 2..Filter3Delka
0x06 | 0x00C0 | 0x00XX | 40192 | List | CHANNE. – CHAN. B – FORM. B | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x00C2 | 0xXXXX | 40194 | Desc | CHANNE. – CHAN. B – DESC. B | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x00C4 | 0x00XX | 40196 | List | CHANNE. – CHAN. B – SWIT. B | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x00C6 | 0x00XX | 40198 | List | CHANNE. – CHAN. B – LOG. B – SAVE B | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x00C8 | 0x000204EEMMMMMM | 40200 | Float | CHANNE. – CHAN. B – LOG. B – FROM B | Lower limit of data storing range | -999..9999
0x10 | 0x00CA | 0x000204EEMMMMMM | 40202 | Float | CHANNE. – CHAN. B – LOG. B – TO B | Upper limit of data storing range | -999..9999
0x10 | 0x00CC | 0x000204EEMMMMMM | 40204 | Float | CHANNE. – CHAN. C – SET. C – MIN C | Scale | -999..9999
0x10 | 0x00CE | 0x000204EEMMMMMM | 40206 | Float | CHANNE. – CHAN. C – SET. C – MAX C | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x00D0 | 0x000204EEMMMMMM | 40208 | Float | CHANNE. – CHAN. C – SET. C – P.TAR. C | Handy set value of tare | -999..9999
0x06 | 0x00D2 | 0x00XX | 40210 | List | CHANNE. – CHAN. C – FILTER – MOD. F.C | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x00D4 | 0x000204EEMMMMMM | 40212 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 0,001..9999
0x10 | 0x00D6 | 0x000204EEMMMMMM | 40214 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 2..Filter2Delka
0x10 | 0x00D8 | 0x000204EEMMMMMM | 40216 | Float | CHANNE. – CHAN. C – FILTER – CON. F.C | Filter constant | 2..Filter3Delka
0x06 | 0x00DA | 0x00XX | 40218 | List | CHANNE. – CHAN. C – FORM. C | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x00DC | 0xXXXX | 40220 | Desc | CHANNE. – CHAN. C – DESC. C | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x00DE | 0x00XX | 40222 | List | CHANNE. – CHAN. C – SWIT. C | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x00E0 | 0x00XX | 40224 | List | CHANNE. – CHAN. C – LOG. C – SAVE C | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x00E2 | 0x000204EEMMMMMM | 40226 | Float | CHANNE. – CHAN. C – LOG. C – FROM C | Lower limit of data storing range | -999..9999
0x10 | 0x00E4 | 0x000204EEMMMMMM | 40228 | Float | CHANNE. – CHAN. C – LOG. C – TO C | Upper limit of data storing range | -999..9999
0x10 | 0x00E6 | 0x000204EEMMMMMM | 40230 | Float | CHANNE. – CHAN. D – SET. D – MIN D | Scale | -999..9999
0x10 | 0x00E8 | 0x000204EEMMMMMM | 40232 | Float | CHANNE. – CHAN. D – SET. D – MAX D | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x00EA | 0x000204EEMMMMMM | 40234 | Float | CHANNE. – CHAN. D – SET. D – P.TAR. D | Handy set value of tare | -999..9999
0x06 | 0x00EC | 0x00XX | 40236 | List | CHANNE. – CHAN. D – FILTER – MOD. F.D | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x00EE | 0x000204EEMMMMMM | 40238 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 0,001..9999
0x10 | 0x00F0 | 0x000204EEMMMMMM | 40240 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 2..Filter2Delka
0x10 | 0x00F2 | 0x000204EEMMMMMM | 40242 | Float | CHANNE. – CHAN. D – FILTER – CON. F.D | Filter constant | 2..Filter3Delka
0x06 | 0x00F4 | 0x00XX | 40244 | List | CHANNE. – CHAN. D – FORM. D | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x00F6 | 0xXXXX | 40246 | Desc | CHANNE. – CHAN. D – DESC. D | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x00F8 | 0x00XX | 40248 | List | CHANNE. – CHAN. D – SWIT. D | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x00FA | 0x00XX | 40250 | List | CHANNE. – CHAN. D – LOG. D – SAVE D | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x00FC | 0x000204EEMMMMMM | 40252 | Float | CHANNE. – CHAN. D – LOG. D – FROM D | Lower limit of data storing range | -999..9999
0x10 | 0x00FE | 0x000204EEMMMMMM | 40254 | Float | CHANNE. – CHAN. D – LOG. D – TO D | Upper limit of data storing range | -999..9999
0x10 | 0x0100 | 0x000204EEMMMMMM | 40256 | Float | CHANNE. – CHAN. E – SET. E – MIN E | Scale | -999..9999
0x10 | 0x0102 | 0x000204EEMMMMMM | 40258 | Float | CHANNE. – CHAN. E – SET. E – MAX E | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x0104 | 0x000204EEMMMMMM | 40260 | Float | CHANNE. – CHAN. E – SET. E – P.TAR. E | Handy set value of tare | -999..9999
0x06 | 0x0106 | 0x00XX | 40262 | List | CHANNE. – CHAN. E – FILTER – MOD. F.E | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x0108 | 0x000204EEMMMMMM | 40264 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 0,001..9999
0x10 | 0x010A | 0x000204EEMMMMMM | 40266 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 2..Filter2Delka
0x10 | 0x010C | 0x000204EEMMMMMM | 40268 | Float | CHANNE. – CHAN. E – FILTER – CON. F.E | Filter constant | 2..Filter3Delka
0x06 | 0x010E | 0x00XX | 40270 | List | CHANNE. – CHAN. E – FORM. E | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x0110 | 0xXXXX | 40272 | Desc | CHANNE. – CHAN. E – DESC. E | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x0112 | 0x00XX | 40274 | List | CHANNE. – CHAN. E – SWIT. E | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x0114 | 0x00XX | 40276 | List | CHANNE. – CHAN. E – LOG. E – SAVE E | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x0116 | 0x000204EEMMMMMM | 40278 | Float | CHANNE. – CHAN. E – LOG. E – FROM E | Lower limit of data storing range | -999..9999
0x10 | 0x0118 | 0x000204EEMMMMMM | 40280 | Float | CHANNE. – CHAN. E – LOG. E – TO E | Upper limit of data storing range | -999..9999
0x10 | 0x011A | 0x000204EEMMMMMM | 40282 | Float | CHANNE. – CHAN. F – SET. F – MIN F | Scale | -999..9999
0x10 | 0x011C | 0x000204EEMMMMMM | 40284 | Float | CHANNE. – CHAN. F – SET. F – MAX F | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x011E | 0x000204EEMMMMMM | 40286 | Float | CHANNE. – CHAN. F – SET. F – P.TAR. F | Handy set value of tare | -999..9999
0x06 | 0x0120 | 0x00XX | 40288 | List | CHANNE. – CHAN. F – FILTER – MOD. F.F | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x0122 | 0x000204EEMMMMMM | 40290 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 0,001..9999
0x10 | 0x0124 | 0x000204EEMMMMMM | 40292 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 2..Filter2Delka
0x10 | 0x0126 | 0x000204EEMMMMMM | 40294 | Float | CHANNE. – CHAN. F – FILTER – CON. F.F | Filter constant | 2..Filter3Delka
0x06 | 0x0128 | 0x00XX | 40296 | List | CHANNE. – CHAN. F – FORM. F | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x012A | 0xXXXX | 40298 | Desc | CHANNE. – CHAN. F – DESC. F | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x012C | 0x00XX | 40300 | List | CHANNE. – CHAN. F – SWIT. F | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x012E | 0x00XX | 40302 | List | CHANNE. – CHAN. F – LOG. F – SAVE F | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x0130 | 0x000204EEMMMMMM | 40304 | Float | CHANNE. – CHAN. F – LOG. F – FROM F | Lower limit of data storing range | -999..9999
0x10 | 0x0132 | 0x000204EEMMMMMM | 40306 | Float | CHANNE. – CHAN. F – LOG. F – TO F | Upper limit of data storing range | -999..9999
0x10 | 0x0134 | 0x000204EEMMMMMM | 40308 | Float | CHANNE. – CHAN. G – SET. G – MIN G | Scale | -999..9999
0x10 | 0x0136 | 0x000204EEMMMMMM | 40310 | Float | CHANNE. – CHAN. G – SET. G – MAX G | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x0138 | 0x000204EEMMMMMM | 40312 | Float | CHANNE. – CHAN. G – SET. G – P.TAR. G | Handy set value of tare | -999..9999
0x06 | 0x013A | 0x00XX | 40314 | List | CHANNE. – CHAN. G – FILTER – MOD. F.G | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x013C | 0x000204EEMMMMMM | 40316 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 0,001..9999
0x10 | 0x013E | 0x000204EEMMMMMM | 40318 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 2..Filter2Delka
0x10 | 0x0140 | 0x000204EEMMMMMM | 40320 | Float | CHANNE. – CHAN. G – FILTER – CON. F.G | Filter constant | 2..Filter3Delka
0x06 | 0x0142 | 0x00XX | 40322 | List | CHANNE. – CHAN. G – FORM. G | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x0144 | 0xXXXX | 40324 | Desc | CHANNE. – CHAN. G – DESC. G | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x0146 | 0x00XX | 40326 | List | CHANNE. – CHAN. G – SWIT. G | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x0148 | 0x00XX | 40328 | List | CHANNE. – CHAN. G – LOG. G – SAVE G | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x014A | 0x000204EEMMMMMM | 40330 | Float | CHANNE. – CHAN. G – LOG. G – FROM G | Lower limit of data storing range | -999..9999
0x10 | 0x014C | 0x000204EEMMMMMM | 40332 | Float | CHANNE. – CHAN. G – LOG. G – TO G | Upper limit of data storing range | -999..9999
0x10 | 0x014E | 0x000204EEMMMMMM | 40334 | Float | CHANNE. – CHAN. H – SET. H – MIN H | Scale | -999..9999
0x10 | 0x0150 | 0x000204EEMMMMMM | 40336 | Float | CHANNE. – CHAN. H – SET. H – MAX H | Value corresponding w/ maximum value of input range | -999..9999
0x10 | 0x0152 | 0x000204EEMMMMMM | 40338 | Float | CHANNE. – CHAN. H – SET. H – P.TAR. H | Handy set value of tare | -999..9999
0x06 | 0x0154 | 0x00XX | 40340 | List | CHANNE. – CHAN. H – FILTER – MOD. F.H | Filter mode | 0=NO, 1=AVER., 2=FLOAT., 3=EXPON., 4=ROUND
0x10 | 0x0156 | 0x000204EEMMMMMM | 40342 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 0,001..9999
0x10 | 0x0158 | 0x000204EEMMMMMM | 40344 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 2..Filter2Delka
0x10 | 0x015A | 0x000204EEMMMMMM | 40346 | Float | CHANNE. – CHAN. H – FILTER – CON. F.H | Filter constant | 2..Filter3Delka
0x06 | 0x015C | 0x00XX | 40348 | List | CHANNE. – CHAN. H – FORM. H | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x015E | 0xXXXX | 40350 | Desc | CHANNE. – CHAN. H – DESC. H | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x0160 | 0x00XX | 40352 | List | CHANNE. – CHAN. H – SWIT. H | Setting of display cycle switching | 0=NO, 1=CHANNE., 2=FILTER, 3=CHAN.+F
0x06 | 0x0162 | 0x00XX | 40354 | List | CHANNE. – CHAN. H – LOG. H – SAVE H | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x0164 | 0x000204EEMMMMMM | 40356 | Float | CHANNE. – CHAN. H – LOG. H – FROM H | Lower limit of data storing range | -999..9999
0x10 | 0x0166 | 0x000204EEMMMMMM | 40358 | Float | CHANNE. – CHAN. H – LOG. H – TO H | Upper limit of data storing range | -999..9999
0x06 | 0x0168 | 0x00XX | 40360 | List | CHANNE. – MAT. FN. – INP. M. | Input value | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x016A | 0x00XX | 40362 | List | CHANNE. – MAT. FN. – MATH. F. | Mathematic function select | 0=MULTIN., 1=1/MUL., 2=LOGAR., 3=EXPON., 4=POWER, 5=ROOT, 6=SIN X, 7=SUMA., 8=DIVIDE., 9=QUADR. .
0x10 | 0x016C | 0x000204EEMMMMMM | 40364 | Float | CHANNE. – MAT. FN. – CON. A | Constant A | -999..9999
0x10 | 0x016E | 0x000204EEMMMMMM | 40366 | Float | CHANNE. – MAT. FN. – CON. B | Constant B | -999..9999
0x10 | 0x0170 | 0x000204EEMMMMMM | 40368 | Float | CHANNE. – MAT. FN. – CON. C | Constant C | -999..9999
0x10 | 0x0172 | 0x000204EEMMMMMM | 40370 | Float | CHANNE. – MAT. FN. – CON. D | Constant D | -999..9999
0x10 | 0x0174 | 0x000204EEMMMMMM | 40372 | Float | CHANNE. – MAT. FN. – CON. E | Constant E | -999..9999
0x10 | 0x0176 | 0x000204EEMMMMMM | 40374 | Float | CHANNE. – MAT. FN. – CON. F | Constant F | -999..9999
0x10 | 0x0178 | 0x000204EEMMMMMM | 40376 | Float | CHANNE. – MAT. FN. – CON. G | Constant G | -999..9999
0x10 | 0x017A | 0x000204EEMMMMMM | 40378 | Float | CHANNE. – MAT. FN. – CON. H | Constant H | -999..9999
0x10 | 0x017C | 0x000204EEMMMMMM | 40380 | Float | CHANNE. – MAT. FN. – CON. I | Constant I | -999..9999
0x10 | 0x017E | 0x000204EEMMMMMM | 40382 | Float | CHANNE. – MAT. FN. – CON. J | Constant J | -999..9999
0x06 | 0x0180 | 0x00XX | 40384 | List | CHANNE. – MAT. FN. – FORM. M | Projection format | 0=000000., 1=00000.o, 2=0000.oo, 3=000.ooo, 4=FLOA. P.
0x06 | 0x0182 | 0xXXXX | 40386 | Desc | CHANNE. – MAT. FN. – DESC. M | Description setting (Measured quantity unit) - 3 ASCII characters | Char(32)..Char(127)
0x06 | 0x0184 | 0x00XX | 40388 | List | CHANNE. – MAT. FN. – SWIT. M | Setting of display cycle switching | 0=NO, 1=YES
0x06 | 0x0186 | 0x00XX | 40390 | List | CHANNE. – MAT. FN. – LOG. M – SAVE M | Channel value storing | 0=NO, 1=ALL, 2=IN, 3=OUT
0x10 | 0x0188 | 0x000204EEMMMMMM | 40392 | Float | CHANNE. – MAT. FN. – LOG. M – FROM M | Lower limit of data storing range | -999..9999
0x10 | 0x018A | 0x000204EEMMMMMM | 40394 | Float | CHANNE. – MAT. FN. – LOG. M – TO M | Upper limit of data storing range | -999..9999
0x06 | 0x018C | 0x00XX | 40396 | List | CHANNE. – MIN.MAX – INP. M.M. | Input value for minimum and maximum | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x018E | 0x00XX | 40398 | CheckBox | CHANNE. – MIN.MAX – CH. M.M. |  | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x0190 | 0x00XX | 40400 | List | OUTPUT. – MEMORY – REWRIT. | Oldest values are rewrited with newest | 0=NO, 1=YES
0x10 | 0x0192 | 0x000204XXXXXXXX | 40402 | Long | OUTPUT. – MEMORY – START | Start of data storing interval | 0..1522491
0x10 | 0x0194 | 0x000204XXXXXXXX | 40404 | Long | OUTPUT. – MEMORY – STOP | End of datastoring interval | 0..1522491
0x10 | 0x0196 | 0x000204XXXXXXXX | 40406 | Long | OUTPUT. – MEMORY – PERIOD. | Period of datastoring | 0..1522491
0x10 | 0x0198 | 0x000204EEMMMMMM | 40408 | Float | OUTPUT. – MEMORY – TRIGG. | Size of pretrigger area of memory | 0..100
0x06 | 0x019A | 0x00XX | 40410 | List | OUTPUT. – LIMITS – LIM 1 – INP. L.1 | Input value for limit 1 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x019C | 0x00XX | 40412 | CheckBox | OUTPUT. – LIMITS – LIM 1 – CH. L.1 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x019E | 0x00XX | 40414 | List | OUTPUT. – LIMITS – LIM 1 – MOD. L.1 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x01A0 | 0x00XX | 40416 | List | OUTPUT. – LIMITS – LIM 1 – TYP. L.1 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x01A2 | 0x000204EEMMMMMM | 40418 | Float | OUTPUT. – LIMITS – LIM 1 – LIM. L.1 | Limit value | -999..9999
0x10 | 0x01A4 | 0x000204EEMMMMMM | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – HYS. L.1 | Hysteresis | 0..9999
0x10 | *0x01A2 | 0x000204EEMMMMMM | 40418 | Float | OUTPUT. – LIMITS – LIM 1 – ON L.1 | Limit value | -999..9999
0x10 | *0x01A4 | 0x000204EEMMMMMM | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – OFF L.1 | Hysteresis | -999..9999
0x10 | *0x01A4 | 0x000204EEMMMMMM | 40420 | Float | OUTPUT. – LIMITS – LIM 1 – PER. L.1 | Hysteresis | 0,001..9999
0x10 | 0x01AC | 0x000204EEMMMMMM | 40428 | Float | OUTPUT. – LIMITS – LIM 1 – TIM. L.1 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x01AE | 0x00XX | 40430 | List | OUTPUT. – LIMITS – LIM 2 – INP. L.2 | Input value for limit 2 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x01B0 | 0x00XX | 40432 | CheckBox | OUTPUT. – LIMITS – LIM 2 – CH. L.2 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x01B2 | 0x00XX | 40434 | List | OUTPUT. – LIMITS – LIM 2 – MOD. L.2 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x01B4 | 0x00XX | 40436 | List | OUTPUT. – LIMITS – LIM 2 – TYP. L.2 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x01B6 | 0x000204EEMMMMMM | 40438 | Float | OUTPUT. – LIMITS – LIM 2 – LIM. L.2 | Limit value | -999..9999
0x10 | 0x01B8 | 0x000204EEMMMMMM | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – HYS. L.2 | Hysteresis | 0..9999
0x10 | *0x01B6 | 0x000204EEMMMMMM | 40438 | Float | OUTPUT. – LIMITS – LIM 2 – ON L.2 | Limit value | -999..9999
0x10 | *0x01B8 | 0x000204EEMMMMMM | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – OFF L.2 | Hysteresis | -999..9999
0x10 | *0x01B8 | 0x000204EEMMMMMM | 40440 | Float | OUTPUT. – LIMITS – LIM 2 – PER. L.2 | Hysteresis | 0,001..9999
0x10 | 0x01C0 | 0x000204EEMMMMMM | 40448 | Float | OUTPUT. – LIMITS – LIM 2 – TIM. L.2 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x01C2 | 0x00XX | 40450 | List | OUTPUT. – LIMITS – LIM 3 – INP. L.3 | Input value for limit 3 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x01C4 | 0x00XX | 40452 | CheckBox | OUTPUT. – LIMITS – LIM 3 – CH. L.3 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x01C6 | 0x00XX | 40454 | List | OUTPUT. – LIMITS – LIM 3 – MOD. L.3 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x01C8 | 0x00XX | 40456 | List | OUTPUT. – LIMITS – LIM 3 – TYP. L.3 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x01CA | 0x000204EEMMMMMM | 40458 | Float | OUTPUT. – LIMITS – LIM 3 – LIM. L.3 | Limit value | -999..9999
0x10 | 0x01CC | 0x000204EEMMMMMM | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – HYS. L.3 | Hysteresis | 0..9999
0x10 | *0x01CA | 0x000204EEMMMMMM | 40458 | Float | OUTPUT. – LIMITS – LIM 3 – ON L.3 | Limit value | -999..9999
0x10 | *0x01CC | 0x000204EEMMMMMM | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – OFF L.3 | Hysteresis | -999..9999
0x10 | *0x01CC | 0x000204EEMMMMMM | 40460 | Float | OUTPUT. – LIMITS – LIM 3 – PER. L.3 | Hysteresis | 0,001..9999
0x10 | 0x01D4 | 0x000204EEMMMMMM | 40468 | Float | OUTPUT. – LIMITS – LIM 3 – TIM. L.3 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x01D6 | 0x00XX | 40470 | List | OUTPUT. – LIMITS – LIM 4 – INP. L.4 | Input value for limit 4 | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x01D8 | 0x00XX | 40472 | CheckBox | OUTPUT. – LIMITS – LIM 4 – CH. L.4 | Each weigh designes channel which will be use. 1 ... channel A, 2 ... B, 4 ... C, 8 ... D, etc. | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x01DA | 0x00XX | 40474 | List | OUTPUT. – LIMITS – LIM 4 – MOD. L.4 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x01DC | 0x00XX | 40476 | List | OUTPUT. – LIMITS – LIM 4 – TYP. L.4 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x01DE | 0x000204EEMMMMMM | 40478 | Float | OUTPUT. – LIMITS – LIM 4 – LIM. L.4 | Limit value | -999..9999
0x10 | 0x01E0 | 0x000204EEMMMMMM | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – HYS. L.4 | Hysteresis | 0..9999
0x10 | *0x01DE | 0x000204EEMMMMMM | 40478 | Float | OUTPUT. – LIMITS – LIM 4 – ON L.4 | Limit value | -999..9999
0x10 | *0x01E0 | 0x000204EEMMMMMM | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – OFF L.4 | Hysteresis | -999..9999
0x10 | *0x01E0 | 0x000204EEMMMMMM | 40480 | Float | OUTPUT. – LIMITS – LIM 4 – PER. L.4 | Hysteresis | 0,001..9999
0x10 | 0x01E8 | 0x000204EEMMMMMM | 40488 | Float | OUTPUT. – LIMITS – LIM 4 – TIM. L.4 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x01EA | 0x00XX | 40490 | List | OUTPUT. – LIMITS – LIM 5 – INP. L.5 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x01EC | 0x00XX | 40492 | CheckBox | OUTPUT. – LIMITS – LIM 5 – CH. L.5 |  | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x01EE | 0x00XX | 40494 | List | OUTPUT. – LIMITS – LIM 5 – MOD. L.5 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x01F0 | 0x00XX | 40496 | List | OUTPUT. – LIMITS – LIM 5 – TYP. L.5 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x01F2 | 0x000204EEMMMMMM | 40498 | Float | OUTPUT. – LIMITS – LIM 5 – LIM. L.5 | Limit value | -999..9999
0x10 | 0x01F4 | 0x000204EEMMMMMM | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – HYS. L.5 | Hysteresis | 0..9999
0x10 | *0x01F2 | 0x000204EEMMMMMM | 40498 | Float | OUTPUT. – LIMITS – LIM 5 – ON L.5 | Limit value | -999..9999
0x10 | *0x01F4 | 0x000204EEMMMMMM | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – OFF L.5 | Hysteresis | -999..9999
0x10 | *0x01F4 | 0x000204EEMMMMMM | 40500 | Float | OUTPUT. – LIMITS – LIM 5 – PER. L.5 | Hysteresis | 0,001..9999
0x10 | 0x01FC | 0x000204EEMMMMMM | 40508 | Float | OUTPUT. – LIMITS – LIM 5 – TIM. L.5. | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x01FE | 0x00XX | 40510 | List | OUTPUT. – LIMITS – LIM 6 – INP. L.6 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0200 | 0x00XX | 40512 | CheckBox | OUTPUT. – LIMITS – LIM 6 – CH. L.6 |  | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x0202 | 0x00XX | 40514 | List | OUTPUT. – LIMITS – LIM 6 – MOD. L.6 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x0204 | 0x00XX | 40516 | List | OUTPUT. – LIMITS – LIM 6 – TYP. L.6 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x0206 | 0x000204EEMMMMMM | 40518 | Float | OUTPUT. – LIMITS – LIM 6 – LIM. L.6 | Limit value | -999..9999
0x10 | 0x0208 | 0x000204EEMMMMMM | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – HYS. L.6 | Hysteresis | 0..9999
0x10 | *0x0206 | 0x000204EEMMMMMM | 40518 | Float | OUTPUT. – LIMITS – LIM 6 – ON L.6 | Limit value | -999..9999
0x10 | *0x0208 | 0x000204EEMMMMMM | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – OFF L.6 | Hysteresis | -999..9999
0x10 | *0x0208 | 0x000204EEMMMMMM | 40520 | Float | OUTPUT. – LIMITS – LIM 6 – PER. L.6 | Hysteresis | 0,001..9999
0x10 | 0x0210 | 0x000204EEMMMMMM | 40528 | Float | OUTPUT. – LIMITS – LIM 6 – TIM. L.6 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x0212 | 0x00XX | 40530 | List | OUTPUT. – LIMITS – LIM 7 – INP. L.7 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0214 | 0x00XX | 40532 | CheckBox | OUTPUT. – LIMITS – LIM 7 – CH. L.7 |  | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x0216 | 0x00XX | 40534 | List | OUTPUT. – LIMITS – LIM 7 – MOD. L.7 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x0218 | 0x00XX | 40536 | List | OUTPUT. – LIMITS – LIM 7 – TYP. L.7 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x021A | 0x000204EEMMMMMM | 40538 | Float | OUTPUT. – LIMITS – LIM 7 – LIM. L.7 | Limit value | -999..9999
0x10 | 0x021C | 0x000204EEMMMMMM | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – HYS. L.7 | Hysteresis | 0..9999
0x10 | *0x021A | 0x000204EEMMMMMM | 40538 | Float | OUTPUT. – LIMITS – LIM 7 – ON L.7 | Limit value | -999..9999
0x10 | *0x021C | 0x000204EEMMMMMM | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – OFF L.7 | Hysteresis | -999..9999
0x10 | *0x021C | 0x000204EEMMMMMM | 40540 | Float | OUTPUT. – LIMITS – LIM 7 – PER. L.7 | Hysteresis | 0,001..9999
0x10 | 0x0224 | 0x000204EEMMMMMM | 40548 | Float | OUTPUT. – LIMITS – LIM 7 – TIM. L.7 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x0226 | 0x00XX | 40550 | List | OUTPUT. – LIMITS – LIM 8 – INP. L.8 | Input value for limit | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0228 | 0x00XX | 40552 | CheckBox | OUTPUT. – LIMITS – LIM 8 – CH. L.8 |  | 0..15(4 channels), 0..255(8 channels)
0x06 | 0x022A | 0x00XX | 40554 | List | OUTPUT. – LIMITS – LIM 8 – MOD. L.8 | Mode of controling limit | 0=HYSTER., 1=FROM. . ., 2=DOSING
0x06 | 0x022C | 0x00XX | 40556 | List | OUTPUT. – LIMITS – LIM 8 – TYP. L.8 | Limit type | 0=CLOSE, 1=OPEN
0x10 | 0x022E | 0x000204EEMMMMMM | 40558 | Float | OUTPUT. – LIMITS – LIM 8 – LIM. L.8 | Limit value | -999..9999
0x10 | 0x0230 | 0x000204EEMMMMMM | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – HYS. L.8 | Hysteresis | 0..9999
0x10 | *0x022E | 0x000204EEMMMMMM | 40558 | Float | OUTPUT. – LIMITS – LIM 8 – ON L.8 | Limit value | -999..9999
0x10 | *0x0230 | 0x000204EEMMMMMM | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – OFF L.8 | Hysteresis | -999..9999
0x10 | *0x0230 | 0x000204EEMMMMMM | 40560 | Float | OUTPUT. – LIMITS – LIM 8 – PER. L.8 | Hysteresis | 0,001..9999
0x10 | 0x0238 | 0x000204EEMMMMMM | 40568 | Float | OUTPUT. – LIMITS – LIM 8 – TIM. L.8 | Time of delay between meeting the condition and the relay switchon/switchoff (according to type of switching) | -99,9..99,9
0x06 | 0x023A | 0x00XX | 40570 | List | OUTPUT. – DATA – BAUD | Transfer rate | 0= 600, 1= 1200, 2= 2400, 3= 4800, 4= 9600, 5= 19200, 6= 38400, 7= 57600, 8=115200, 9=230400
0x10 | 0x023C | 0x000204EEMMMMMM | 40572 | Float | OUTPUT. – DATA – ADDR. | Instrument address | 0..31
0x10 | 0x0240 | 0x000204EEMMMMMM | 40576 | Float | OUTPUT. – DATA – ADR. P.B. | Instrument address on the field bus | 0..125
0x06 | 0x0242 | 0x00XX | 40578 | List | OUTPUT. – DATA – PROT. | Communication protocol | 0=ASCII, 1=M. BUS, 2=MODBUS
0x06 | 0x0244 | 0x00XX | 40580 | List | OUTPUT. – AN. OUT. – INP. A.O. | Input value for analog output | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x0246 | 0x00XX | 40582 | List | OUTPUT. – AN. OUT. – TYP. A.O. | Type | 0=0-20mA, 1=Er4- T, 2=4-20 T, 3=Er4-20, 4=4-20mA, 5=0- 5mA, 6=0- 2 V, 7=0- 5 V, 8=0-10 V, 9=+-10 V
0x10 | 0x0248 | 0x000204EEMMMMMM | 40584 | Float | OUTPUT. – AN. OUT. – MIN A.O. | Minimum display value corresponding with bottom value of selected A.O. range | -999..9999
0x10 | 0x024A | 0x000204EEMMMMMM | 40586 | Float | OUTPUT. – AN. OUT. – MAX A.O. | Maximum display value corresponding with top value of selected A.O. range | -999..9999
0x06 | 0x024C | 0x00XX | 40588 | List | OUTPUT. – DISP. – PERM. | Value to be permanent displayed | 0=NO, 1=CHAN. A, 2=FIL. A, 3=CHAN. B, 4=FIL. B, 5=CHAN. C, 6=FIL. C, 7=CHAN. D, 8=FIL. D, 9=CHAN. E, 10=FIL. E, 11=CHAN. F, 12=FIL. F, 13=CHAN. G, 14=FIL. G, 15=CHAN. H, 16=FIL. H, 17=MAT. FN., 18=MIN., 19=MAX., 20=LIM. 1, 21=LIM. 2, 22=LIM. 3, 23=LIM. 4, 24=LIM. 5, 25=LIM. 6, 26=LIM. 7, 27=LIM. 8, 28=TIME, 29=DATE, 30=TARE A, 31=TARE B, 32=TARE C, 33=TARE D, 34=TARE E, 35=TARE F, 36=TARE G, 37=TARE H, 38=TAR.ACT, 39=P. T. A, 40=P. T. B, 41=P. T. C, 42=P. T. D, 43=P. T. E, 44=P. T. F, 45=P. T. G, 46=P. T. H, 47=P.T. ACT., 48=COLD. J., 49=ALL. CH., 50=ALL. F., 51=DISPL.
0x06 | 0x024E | 0x00XX | 40590 | List | OUTPUT. – DISP. – CHAN. | This item intend actual channel number showing. | 0=NO, 1=NUMBER., 2=CHARS.
0x06 | 0x0250 | 0x00XX | 40592 | List | OUTPUT. – DISP. – BRIGHT | Brightness of displey | 0= 0%, 1= 25%, 2= 50%, 3= 75%, 4= 100%
0x06 | 0x0260 | 0x00XX | 40608 | List | SERVIC. – LANG. | Language variation of menu | 0=CZECH, 1=ENGL.